## BoQ-PRC IQD
Annex B - BoQ - Medical Devices for PRC in Telafer General Hospital
Notes: • The images shown are for illustration purposes only and may not be an exact representation of the product. • All medical devices should be provided with a guarantee letter (between 2 - 5 years or what time suits the item's nature and functionality lifetime from the supplier company or contractor and to be confirmed with its validation with the relevant authorities.
No. | Description | QTY | Origin | Unit | Unit price | Sub-total | Illustration Photo
1 | Training stairs (step climber) adults + children: Made with a solid hardwood/plywood base for strength and durability. Non-Slip Straight Training Stairs with Bus Step. Training Stairs, 36 in. Wide, Straight. Fixed with Sturdy handrails . Non-Slip surface to minimize the risk of falling. Wieght capacity up to 181 Kg. Enclosed on Two Sides. Overall Dimensions 116" L x 36". Height to Platform24". Number of Steps 8 &4. Riser Size 3"x 10"& 6''x10''. Platform Size 36" x 36". | 2 |  | Unit |  | 0
2 | Patient examination Treatment Table: Examination table, 2 sections. • Mounted on 4 sturdy supports, all finished with height-adjustable feet. • Both sections are fitted with non-removable padded upholstery. • Backrest angle adjustable via secured pawl and gear ratchet, safe for patient and operator. • When fully extended, both sections align to a perfectly flat surface. • Transfer bars connect all lower distal portions of the 4 supports, providing maximal structural strength. Materials: • High resistance to corrosion (tropical environment) • Frame: epoxy-coated tubular steel • Adjustable feet: rubber or nylon • Padded upholstery: high-density polyurethane foam, density 28-30 kg/m3 • Cover: plastic, flexible, highly tear resistant, anti-static, flame-retardant, non-absorbing, waterproof, and cleanable with hospital-grade disinfection products. Dimensions: • Examination table, two sections extended, including upholstery: (185-190) x 55 x 80cm (l x w x h). • Frame: 3 cm (outside, across), 1.35-1.5 mm thickness. • Upholstery: 4.5-5 cm (h). • Carrying capacity: minimum 160 kg. | 6 |  | Unit |  | 0
3 | Electric medical Treadmill (150 - 250 KG) Application: physiotherapy and rehabilitation centers. Monitor YES , LCD or better. Parameters monitoredL: Time, speed, elevation, distance, pulse, calories Upgradable software. Dimensions (l x w ) ≥ (210 x 65 cm). Running Area (l x w ) ≥ (150 x 50 cm). Deck :Shock load reduction, for the joint. Belt surface with non slip material. Motor: DC motor, ≥ 3 HP. Speed range: Forward : 0-12 km/hr. or more. Optional Reverse: 0-4.8 km/hr. or more. Elevation: 0-±10% grade or more. Patient capacity: 180 kg or more. Side bars: YES adjustable. Emergency stop: YES. Training programmes: YES ≥10. User-defined programmes : YES ≥ 15. Options : Additional side bars. Ramp ,Power & Environmental Requirements. Power input to be 220-240VAC, 50Hz. The supplier should be confirmed that: The Equipment is suitable for work in climate conditions of Iraq in terms of temperature and humidity. | 2 |  | Unit |  | 0
4 | Exercise bike: Application: cardio-respiratory training, rehabilitation and sports training, exercising. of patients with movement impairments and disabilities. Frame Steel or better. Breaks mechanical brake system. Seat Adjustments Vertical adjustment. crank adjustable. Read Outs Distance, Speed, Pulse. Max. patient weight ≥ 180 kg. Load 20-400 W or a wider range. Input power 220/240 VAC, 50/60 Hz single phase. Three Pin type Plug, Voltage stabilizing and over current protection circuit. Environmental requirements The unit shall be capable of being stored and operated in an Iraq environment. | 3 |  | Unit |  | 0
5 | Parallel bar for adult and children 2m: Dimensions (cm):211 x 81 x 87 H. Handrail height adjustment system:double gas sprint. Min – max handrail height (cm):64 - 80. Handrail width adjustment system:Manual, 3 positions. Handrail width min - max (cm):55 -71. Individual handrail loadbearing capacity:135 Kg. Modularity (combination of more parallel bars):Removable access slide. | 2 |  | Unit |  | 0
6 | Parallel bar for adult and children 3m: Max dimensions (cm):311 x 81 x 108 H. Handrail height adjustment system:double gas sprint. Min – max handrail height (cm):74 - 102. Handrail width adjustment system:Manual, 3 positions. Handrail width min - max (cm):55 - 71 Individual handrail loadbearing capacity:135 Kg. Modularity (combination of more parallel bars): Removable access slide. | 2 |  | Unit |  | 0
7 | Walking frame adult: Type: Fixed Material: Aluminium Size: Adult Overall Width (mm): 600 Handle Height (mm):740 Tip Size (mm): 25.4 Max User Weight (kg): 200 | 4 |  | Unit |  | 0
8 | Childrens Wheeled Walking Frame: One Childrens Wheeled Walking Frame. Strong aluminium frame with non-marking tips. Height adjustable for a tailored fit. Soft contoured handles for comfort. Ferrule size: 25mm (2.5cm). Maximum user weight: 16st (100kgs). Height: 55cm-66cm. | 2 |  | Unit |  | 0
9 | Wax bath device (Including wax): Application: used for storing and melting wax at desired temperature level for physiotherapy application. FDA CLEARANCE or CE MARK: Yes ( FDA Preferred). Tank Material stainless steel. Range of thermostat 30 – 90º C or wider range. Tank capacity ≥ 17 Liter. Alert indicator End of run, over/under temperature. Hot packs tank Optional. Time control Yes. Safety Automatic switch-off when overheated, door safety. Trolley Yes. Accessories ≥10 Kg. Paraffin Wax, brush for wax application. Power & Environmental Requirements, Power input to be 220-240VAC, 50Hz. The supplier should be confirmed that: The Equipment is suitable for work in the climate conditions of Iraq in terms of temperature and humidity. | 2 |  | Unit |  | 0
10 | Infrared (IR) device Power 250w or more: Configuration/ on casters Light source Infrared lamp protected with shiny polished aluminum Cable length ≥ 3 m Treatment timer 1 – 30 minutes or wider range STANDARD ACCESSORIES protective goggles for patient protective goggles for therapist Radiator 250w or more Power & Environmental Requirements • Power input to be 220-240VAC, 50Hz. The supplier should be confirmed that: The Equipment is suitable for work in climate conditions of Iraq in terms of temperature and humidity. | 2 |  | Unit |  | 0
11 | shockwave thermal therapy device : Application Used for but not limited to (dissolving of calcified fibroblasts, increase of collagen production, decrease of muscle tension, analgesic effect, and muscle relaxation, Tendon problems, wound healing ). Technology: Electromagnetic or Piezoelectric Features: LCD color display Touch-Screen or external I pad ≥ 20 treatment programs or treatment protocols. Frequency: (1-≥ 8) Hz. Energy Density: (0.01 – ≥ 0.5) mj/mm². Penetration Length: (0 - ≥ 20) mm. Lifetime:.≥ 1 million shock. Modes: Single mode, continuous mode, others to be specified. Applicator heads: ≥2 (different sizes and penetration length). Trolley: yes Upgradable: Yes. Power supply: Voltage: 220~240 VAC, 50/60 Hz three pin G- type plug. Other specification: Interface to PC or other Accessories: Gel 1000ml and operating manual/ 2pcs. Holder for hand-piece. Silicon covers. Mains cable. Foot switch control or handpiece control. | 1 |  | Unit |  | 0
12 | X-ray LED View Box: Supply, install and test X-Ray LED View Box with Automatic Film Activation and Variable Brightness Control, Color: White Holding: Plastic Clips WITH SENSOR for X-ray film Holding. Handling: Light Weight & Portable Approx. Viewing Area : 71 x 35.5 cm. | 4 |  | Unit |  | 0
13 | Trolly Size 60*40*80cm: Smooth stainless steel shelves with handrail on three sides. Two stainless steel shelves with three sided railing on both shelf. Four 75mm silent castors and two with brakes are equipped. The guardrails are equipped to prevent objects from falling. Polished and welded well. Firm and stable installation. Full Stainless Steel, Four caster wheels (2 Lockable) | 4 |  | Unit |  | 0
14 | Multiple exercise device (arm, quadriceps, shoulder): Application: used for general treatment techniques in rehabilitation centers, Neurological exercise therapy. Type hydraulic without castors. Sections ≥ 3. Frame Rigid design Treatment surface (L*W): ≥ (189*75 cm) Head or Back section Adjustable Lifting capacity ≥ 200 kg Height Adjustable (45-95) cm or wider range UPHOLSTERY Mattress: flame retardant, waterproof FOAM covered with leather Nose-hole YES Accessories : Shoulder/neck support FIXATION STRAPS ≥ 3 ROLLS AND CUSHIONS ≥ 10 different sizes & shapes . Input power: 220/240 VAC, 50/60 Hz single phase. Three Pin type Plug. | 1 |  | Unit |  | 0
15 | Traction (lumber + cervical)Traction Device : Therapy modes continuous, intermittent, progressive (or compatible.) FDA CLEARANCE or CE MARK Yes ( FDA Preferred). Control YES, Microcomputer. Display digital / LCD / touchscreen. Traction force (1.5 – 60) kg or wider range. Treatment time According to treatment programs. Parameter memory ≥ 20 cases. Safety control. for traction force YES. Traction Table: Type Electric Or Hydraulic. Table’s section ≥ 2. Adjustable height 45-95 cm or wider range. Head section Adjustable. Mattress flame retardant, water proof FOAM covered with leather. Lifting capacity ≥ 170 kg. nose-hole YES. Footswitch height adjustment YES. | 1 |  | Unit |  | 0
16 | TENS Galvanic device: (Transcutaneous Electrical Nerve Stimulation (TENS)) COMBO Mini (TENS & Muscle Stimulator) with Two Channel Tens . A mini Portable muscle stimulation unit that delivers Tens, Interrupted Galvanic and Surged Faradic currents for therapeutic use. | 4 |  | Unit |  | 0
17 | Standing device for children: Recommended for children, and young people in the following cases: Infantile cerebral palsy - different forms, Meningomyelocele, Brain and spinal cord injuries. For children below the age of 10, for children height 100 - 130 cm . Max user weight 60 Kg | 2 |  | Unit |  | 0
18 | Sitting device for children: Provides head support, spinal alignment, and pelvic stability (balance) to children with physical disabilities. Special seating also prevents muscle contraction and fatigue. | 2 |  | Unit |  | 0
19 | Balances Balls with 45cm, 55cm, 65cm, and 75cm in diameters. Material: PVC, thick. The accessories are 1 pump, 1 air puller, and 2 air plugs. | 4 |  | Unit |  | 0
20 | Children's Pillow Set : Bolsters, wedges, or pillows the quality of the material is firm, high-density urethane foam core. Along with a knitted back, sewn with a vinyl cover. Small Pillow (14″ x 12″ x 3″). Large Pillow (13″ x 20″ x 3″). Round 6″ Bolster (6″ x 0″ x 24″). Round 8″ Bolster (8″ x 0″ x 24″). Half-cut Bolster (18″ x 8″ x 6″). Flat Cube Pillow (12″ x 8″ x 3″). Cube (12″ x 8″ x 6″). Wedge Pillow (18″ x 24″ x 6″). Wedge Pillow (32″ x 20″ x 12″). | 2 |  | Unit |  | 0
21 | Cando Accordion Fold Mat: The CanDo Accordion Fold Mats fold every two feet along their length for easy, compact storage and convenient portability, and come with Velcro fasteners at the ends that can be attached for a larger mat surface | 2 |  | unit |  | 0
22 | Exercise Therapy Balance Board: Material Wood. Item Dimensions LxWxH 17.9 x 13.95 x 3.5 inches. Maximum Weight Recommendation 350 Pounds. L 56 cm , W 35 cm, H 12 cm. | 2 |  | Unit |  | 0
23 | Balance Trainer: Material: Rubbe.r Style: Inflatable. Latex-free, burst-resistant material. 6 rubberized feet. Holds up to 136 kg (300 lbs.). Weight: 9 kg (13 lbs.). Diameter: 65 cm. Height: 25.4 cm (10”) when fully inflated. Includes hand pump and owner’s manual. | 2 |  | Unit |  | 0
24 | Crawling devices for below 2 years of age children mainly: With Spina Bifida Deal for children up to 50 lbs. Featherweight at only 7lbs Provides support to patient while instructing, encouraging and reinforcing crawling. The aluminum frame and foam-padded, suspended harness are height adjustable. The caregiver can steer from the rear to provide necessary assistance. Accommodates 24 Kg. | 2 |  | Unit |  | 0
25 | Wall bar Stainless steel: Wall Bars / Stall Bars / Swedish ladder: 13 Oval bars made from wood, and have 40 x 30 mm round rung.s The mounting of bars to both sides of the product is performed by using wood screws on the back side. Space between the rungs 17.5 cm, oval rungs except that the first rung to be offset by about 21.5 cm; and the distance between the bars and the wall to be about 12.5 cm. Stall bars can be installed on stud framing with the supplied slats. Drilling for proper framing location is required. Weight capacity 150 kg (it mean that every single horizontal rod can support this weight). Dimensions: 230 x 90 cm. | 2 |  | Unit |  | 0
26 | Massage device Application: Reduces of pain and soothes aches, suitable for shoulder, back, hips, upper legs, and calf. Handheld. Lightweight. Variable speeds ≥2. Frequency : (1200 – 3000) Powerful massage motor (100 – 250) VA. (2 – 3) m power cord. Changeable massager heads. Option: - infrared LED lamp (built-in). Power and environmental requirements: Power input to be 220-240VAC, 50Hz. The supplier should be confirmed that: The Equipment is suitable for work in the climate conditions of Iraq in terms of temperature and humidity. | 4 |  | Unit |  | 0
27 | IRIS Balance Pad : Foam Large Yoga Mat Trainer for Physical Therapy, Stability Workout, Knee and Ankle Exercise, Strength Training, Rehab - Chair Cushion for Adults, Kids. Default size. | 10 |  | Unit |  | 0
28 | Swiss ball Unisex: Anti-Burst & anti-slip Gym Ball with Foot Pump ¦ Exercise Rubber Ball ¦ Anti Burst Swiss Birthing Stability Ball for Workout, Pregnancy, Balance, (Multicolor, Sizes: 55 to 85 cm). | 4 |  | Unit |  | 0
29 | Three movement Electric Treatment Couch : Safe working weight capacity of 230kg (36 stone). Height range from 46cm (18.1”) to 91cm (35.8”). 1 motorized actuator for height adjustment. Gas-assisted head section from -35° to 87°. Gas-assisted foot section from -29° to 20°. 4 retractable heavy-duty 75mm castors . Twin oversized pillars for added stability. Antimicrobial upholstery. Foot control fitted as standard. Hand control available on request. Range of matching accessories available. Breathing hole as standard. | 2 |  | Unit |  | 0
30 | Wheelchair Medium size: Foldable Wheelchair with Sturdy X-frame Easy Folding Mechanism, Steel Chassis, and Rear Mag Wheels. Net weight of the wheelchai.r 19.9kg/43.8lb (medium). | 2 |  | Unit |  | 0
31 | Wheelchair Large size: Foldable Wheelchair with Sturdy X-frame Easy Folding Mechanism, Steel Chassis, and Rear Mag Wheels. Net weight of the wheelchair. 20.9kg/46lb (large). | 2 |  | unit |  | 0
32 | Medical Privacy Screen: • Hospital bedscreen, 3 sections with curtains. • Bedscreen mounted on swivel castors; positioned in broad stance for good stability. • Curtains integrated centrally in the frame, leave no gaps near the hinges. • Material: • High resistance to corrosion (tropical environment). • Frame: Epoxy or enamel coated or chrome plated tubular steel. • Curtains: Plastic, non-transparent, tear resistant, flame retardant, disinfectant- and liquid proof, • washable. Light colored. • Caster frame/bracket: steel or nylon Dimensions: Bedscreen: 3 panels of approx. (60-65) x (195-200) cm (w x h) each. Swivel castors: Approx. diameter 5 cm. Knockdown construction: yes Supplied with: 1 x complete set of tools required for assembly. 1 x set of plastic curtains. List of accessories and parts. Detailed step-by-step instructions for assembly and safe use, text-free pictorial based (i.e. line-drawings only). | 10 |  | unit |  | 0
33 | Mobile Posture Mirror: Or patient assessments, postural training in various rehabilitation clinicsز A steel frame with powder coating for durability where the mirror pivots centrally. Moves on four hard wearing plastic castors. 4 non marking castors. Safety backed mirror glass. Can be tilted to give the best possible view. Mirror size: 1200mm x 360mm. | 4 |  | unit |  | 0
| Total IQD |  |  |  |  | 0

## BoQ-Intisar PHCC IQD
Annex B - Medical Devices for PHCC in Telafer General Hospital
Notes: • The images shown are for illustration purposes only and may not be an exact representation of the product. • All medical devices should be provided with a guarantee letter (between 2 - 5 years or what time suits the item's nature and functionality lifetime from the supplier company or contractor and to be confirmed with its validation with the relevant authorities.
No. | Device Name اسم الجهاز | Details تفاصيل الجهاز | Qty | Origin | Unit | Illustration Image | Unit Price | Sub-total
1 | LAB MICROSCOPE مجهر ضوئي | MICROSCOPE SPECIFICATION: • Purpose: to provide magnification and resolution of specimens under examination. • Heavy duty and compatible design. • OBSERVATION TUBES Type: binocular. • Eyepieces: 4x, 10x, 40x, 100x, wide field, high-eye point. • Interpapillary distance adjustment; mm: 50-75. • NOSEPIECE CONFIGURATION: 4-7 positions. • Objectives Magnification: Achromatic, plan achromatic, phase contrast. • TOTAL MAGNIFICATION: 20x-1,500x. • STAND Focusing mechanism: Moving stage, rack, and pinion. • Coarse, fine adjustments: Coaxial, low mounted. STAGE Type: Mechanical. Tension adjustment:Coaxial. • Stage Motion: X-Y. Condenser type: focusable. Numerical aperture: 0.9 to 1.4. • Light source: halogen lamp (light blue without heat) with replace- able socket, external power supply. • CONTRAST METHODS: Bright field, phase, dark field, polarization, fluorescence • POWER SUPPLY:220-740V 50-60Hz | 1 |  | Unit |  |  | 0
2 | BENCHTOP CENTRIFUGE جهاز الطرد المركزي | • Spins support 3ml-15ml test tubes at rpm > 5000. • No. of tubes: 32 tubes. • Max RPM: between (5.000 - 10,000) rpm. • Digital programmable control and readout. • high g-forces for fast separation • Variable speed and time settings. • Lid safety lock. • Rubber feet to prevent slipping • impact-resistant lid • Wide mouth for easy clean-up • Metal rotor and tube shields • Zero-rpm interlocking lid • Brushless motor • Programmable time and scoped • Angled and (swinging) tube holder as an alternative • Fast breaking • POWER REQUIRMENTS:220v ,50 Hz * Noise level should be less than 70 dB. * Audible alarm, imbalance & end of run. | 1 |  | Unit |  |  | 0
3 | Spectrophotometer جهاز المطياف الضوئي | Used in the quantitative determination of solutions Heavy duty and compatible design Microprocessor-based, grating type single beam Spectrophotometer with built-in data logging in. The unit should have a self-diagnostics compartment with a built-in 16 mm test-tube holder and can accept round and square tubes. • Wavelength Range: 380 - 800 nm • Spectral Bandwidth: 5nm • Wavelength Accuracy: 12± nm, • Wavelength resolution: 0.1 nm. • Lamp type: Quartz halogen or tungsten lamp. • Photometric Range: absorbance ( -2.5 to 2.5) Abs & transmittance 1% to 200%T. • Tube size: round tube and cuvette and flow cell (oreferred\Photoretric linearitv: <1% for A<2.000 in the range 340 to 900nm. • Photometric Accuracy (in the range for 200 to 1000nm): 0.003 AA<0.600,0.5% of the reading for 0.600 <A< 2.000. • Automatic cell reorganization facility, • Self-Test and Auto Calibration Modes. • Methods and profiles: $100 user defined methods. 20 profiles each for kinetics and spectra. • Detector: Silicon photodiode. • Output: LCD display with capability to interface with any PC. • Power line: 220-240VAC:50-60Hz with online UPS: built in printer (preferred). | 1 |  | Unit |  |  | 0
#REF!
4 | Autoclave جهاز اوتوكليف | • Sterilization of surgical solid instruments, hollow instruments. • Heavy duty & compatible design. • Microprocessor controlled. • Self-evaluations on the process of sterilization, and automatic failure detection report. • Instant vaporizer outside the chamber. • Chamber: (18~23) liter, stainless steel. • Integrated printer (optional). • Robot-controlled locking system. • LCD display. • Water loading pipe with filter and direct draining to the water piping system. • Possibility of serial connection to a PC, with cycle data recording. • Tray holder contains 22 trays. • Vacuum pump performing two stages (pre & post vacuum). • Thermal drying system. • Full water alarm system. • Input power: 220/240 VAC, 50/60 Hz single phase. • Accessories: tray extraction handle. | 1 |  | Unit |  |  | 0
5 | Blood Glucose Meters جهاز فحص السكر بالدم | Application: measure blood glucose concentration using a reagent strip and a drop of capillary blood from a finger puncture that clinical use (operating theater). with 1000 test. • Patient type: adult, pediatric, neonate. • Methodology: electrochemical or reflectance photometry(preferred) and other. • Enzyme(s) used: Glucose oxidase or Glucose dehydrogenase or other reagent. • Strips: yes. • Specimen type: capillary whole blood, venous blood, arterial blood, neonatal. blood; minimum sample volume, ≥ 0.3µI. • Measurement range, mg/di: 10- 600. • Test time, sec: between 3 to 60. • Meter calibration checks: yes. • Lot-specific calibration: yes. • Readout: LCD. • Alert indicators: high/low levels: out of measuring range, out of temperature range hypoglycemia, hyperglycemia, low battery, under-dose detection. • Automatic shutoff after 2 min. • Data management :yes. • Number of tests ≥ 250. • Date/time stored with test results: yes. • Life, no. of tests ≥ 1,000 • Battery: yes environmental conditions the equipment suitable for working in the climate conditions of Iraq in terms of temperature & humidity. | 4 |  | Unit |  |  | 0
#REF!
6 | Micropipette الماصة المجهرية 1-100 ul | • Variable Pipetting volumes 10-100ul • Smooth, Polypropylene material, prevents liquid retention. • Inert self-sealing barrier filter for prevention of aerosol contamination of pipette and sample • Sterile & free of DNA, DNase, RNase, and endotoxins. • Racked primary packaging, 96 pipette tips per rack. • Universal fit, compatible with all standard micropipettes. • Shelf life: 24-36 months - Weight & volume: Estimated Weight: 1.58kg Estimated Volume: 9.5dm3 | 1 |  | Unit |  |  | 0
7 | Micropipette الماصة المجهرية 100-1000ul | Variable Pipetting volumes 100-1000ul ·Smooth, Polypropylene material, prevents liquid retention ·Racked primary packaging, 96 pipette tips per rack ·Universal fit, compatible with all standard micropipettes Weight & volume: Estimated Weight: 1.96kg Estimated Volume: 13.10dm3 | 1 |  | Unit |  |  | 0
8 | Nebulizer جهاز تبخير | Nebuliser for the treatment of upper/lower respiratory system and for the delivery of nebulised drugs. • The unit is compressor driven • Depending on the model supplied: either portable or a handheld model. • If equipped with a compressor, the compressor is maintenance and oil-free. • Air inlet filter (depending on the model supplied). • Depending on the model supplied: either continuous operation or auto switch off after 20 minutes but allows for immediately restart. • For compressor model: air flow at least 5 L/min. • Minimum drug nebulisation flow greater than: 0.2 ml/min. • Integrated drug reservoir with a minimum capacity: 6 ml. • Residual volume (not operating below) less than 2 ml. • Median Mass Aerodynamic Diameter (MMAD) of particles 2.44 µm. • Noise level ≤ 55 dB • Sturdy construction, suitable to be disinfected with hospital-grade products. • Power requirements: 240 Volts - 50/60 Hz (110 Volt available, specify when ordering). SUPPLIED WITH Instructions for assembly, use and maintenance in English. +1 x carry bag (depending on model supplied). 2 x reusable mouthpiece.+2 x air tubing set.+4 x reusable adult face masks.+4 x reusable paediatric face masks. 5 x air filters (depending on the model supplied).+2 x nebuliser tank. | 2 |  | Unit |  |  | 0
#REF!
9 | Sphygmomanometer جهاز قياس الضغط | • Composed of a cuff containing an inflatable bag. • The inflatable bag is connected via a tube to a bulb with an integrated manometer needle gauge. • The cuff is made of durable material (e.g. nylon), which is non-deformable, and washable at 30°C. • The cuff is fitted with double Velcro fastening, enabling a tight and secure fit around the arms. • The cuff is reinforced on both sides. • Size cuff for adults. • The bag is inflated by means of a flexible bulb connected via a tube. • Material tube: rubber or other suitable material, e.g. silicone rubber, crack resistant. • Length tube between 50 to 70 cm. • Gauge graduated 0 - 300mmHg (min) in 2 (max) mmHg increments, with pressure release valve. • Mercury design. | 4 |  | Unit |  |  | 0
10 | Snellen Chart Label لوحة فحص البصر | Working distance: 20 feet - 6.1 m • Size: 23 x 35.5 cm Polypropylene natural white sheet, approx. thickness 0.8 to 1.0 mm | 1 |  | Unit |  |  | 0
11 | Portable Stand lamp بلاجكتور فحص محمول | • Heavy base with low centre of gravity. • Single light head with LED lights in cluster. • Articulating arm between light and base stand. • Base with 5 anti-static swivel castors in a star formation. • At least two castors have been equipped with brakes. • Minimum 14 cm fixed focus field of view. • Minimum illumination of 60,000 lux at 0.5m. • Minimum colour temperature between 4,000°K. • Minimum colour rendering index (CRI) of 95. • Minimum LED life more than 40,000 hours. • Floor to light height adjustable to include at least range of 1.1 m to 1.75 m. • On/off switch. • Power requirements: 100 - 240 Volts - 50/60 Hz. | 2 |  | Unit |  |  | 0
12 | Patient privacy curtain برافان | • Hospital bedscreen, 3 sections with curtains. • Bedscreen mounted on swivel castors; positioned in broad stance for good stability. • Curtains integrated centrally in the frame, leave no gaps near the hinges. • Material: • High resistance to corrosion (tropical environment). • Frame: Epoxy or enamel coated or chrome plated tubular steel. • Curtains: Plastic, non-transparent, tear resistant, flame retardant, disinfectant- and liquid proof, • washable. Light colored. • Caster frame/bracket: steel or nylon Dimensions: Bedscreen: 3 panels of approx. (60-65) x (195-200) cm (w x h) each. Swivel castors: Approx. diameter 5 cm. Knockdown construction: yes Supplied with: 1 x complete set of tools required for assembly. 1 x set of plastic curtains. List of accessories and parts. Detailed step-by-step instructions for assembly and safe use, text-free pictorial based (i.e. line-drawings only). | 3 |  | Unit |  |  | 0
#REF!
13 | Examination bed سرير فحص درجة أولى مع سلم للصعود | Examination table, 2 sections. • Mounted on 4 sturdy supports, all finished with height-adjustable feet. • Both sections are fitted with non-removable padded upholstery. • Backrest angle adjustable via secured pawl and gear ratchet, safe for patient and operator. • When fully extended, both sections align to a perfectly flat surface. • Transfer bars connect all lower distal portions of the 4 supports, providing maximal structural strength. Materials: • High resistance to corrosion (tropical environment) • Frame: epoxy-coated tubular steel • Adjustable feet: rubber or nylon • Padded upholstery: high-density polyurethane foam, density 28-30 kg/m3 • Cover: plastic, flexible, highly tear resistant, anti-static, flame-retardant, non-absorbing, waterproof, and cleanable with hospital-grade disinfection products. Dimensions: • Examination table, two sections extended, including upholstery: (185-190) x 55 x 80cm (l x w x h). • Frame: 3 cm (outside, across), 1.35-1.5 mm thickness. • Upholstery: 4.5-5 cm (h). • Carrying capacity: minimum 160 kg. | 4 |  | Unit |  |  | 0
14 | Adult Scale ميزان للكبار | Weighing Scale for Adults with height measuring rod. Steel body duly epoxy powder coated. Heavy standing base covered with Antiskid mat. Capacity 180 kg x 100g graduation. Height Rod 75-200cm x 5mm graduation Specially designed for hospitals and health centres.... | 1 |  | Unit |  |  | 0
15 | Baby Scale ميزان طفل | • Weighing range 0 - 10kg • Minimum graduation: 5g. • Accuracy: ±5g. • Display in kilograms. • Functions: automatic switch-off, TARE, HOLD, auto-calibration with each switch-on. • Large LCD display. • Reading time max 5 seconds. • Splashproof and shock-resistant lightweight body. • Smooth surface/finishing for easy cleaning/disinfection. • All vital parts are made of rust-proof, child-friendly materials; display covers are made of durable, transparent plastic. • Color: white or similar: off-white, cream, light grey, etc. • Power requirements: 220 V/50Hz. • Additional power supply: customer-replaceable batteries. | 1 |  | Unit |  |  | 0
16 | Intravenous Hock Stand | • Movable height adjustable infusion stand. • Heavy carriage mounted on 5 swivel castors. • The lower-end support column is deep and securely fixed into the carriage base. • A solid manual lever at the upper end of the support column, allows setting telescopic upper part at the required height. • A brake prevents exceeding the maximum height setting. • Double hook fixed at the top of the telescopic rod. Material: • High resistance to corrosion (tropical environment), support, and telescopic column: Austenitic stainless steel grade 304. • Carriage-base: single molded unit, polypropylene, castor frame/bracket: polypropylene or nylon, Wheel bearing: polypropylene or nylon. • Castor fixation into carriage base: stainless steel. Dimensions: • Height, adjustable: 135-225 cm, carriage base, diameter: 55-58 cm, support column: 3 cm (outside, across), 1-1.5mm thickness. • Telescopic upper part: 2.5 cm (outside, across), 1-1.5mm (thickness), swivel castor twin-wheel: (4-5) x 5 cm (w x diameter). • Carrying capacity: 10 kg, knockdown construction: yes. Items supplied with: • 1 x complete set of tools required for assembly, list of accessories and parts Detailed step-by-step instructions for assembly and safe use, text-free pictorial based (i.e. line-drawings only) | 3 |  | Unit |  |  | 0
#REF!
Total IQD